## Zircon U-Pb + trace elements
Table S5. LA-ICP-MS U-Pb isotopic and trace element analyses on zircon - Metadata
Laboratory and Sample Preparation
Laboratory name | Dept. of Earth Science, Institute for Geochemistry and Petrology - ETH Zurich
Sample type/mineral | Natural zircon
Sample preparation | Conventional mineral separation, 1 inch resin mount, polishing with diamond suspension
Imaging | Yes
Laser ablation system
Make, Model and type | ASI (Resonetics) RESOolution S-155
Ablation cell and volume | Laurin Technic, dual-volume S-155 ablation cell, effective volume ca. 1 cm3
Laser wavelength (nm) | 193 nm
Pulse width (ns) | 25 ns
Fluence (J cm-2) | ca. 2.5 J cm-2
Repetition rate (Hz) | 3 Hz
Spot diameter (mm) | 29 mm
Ablation duration | 25 s
Ablation pit depth / ablation rate | ca. 6 mm pit depth, equivalent to average 0.08 μm/pulse
Sampling mode / pattern | Static spot ablation
Carrier gas | 100% He in the cell, Ar (+ N2) make-up gas combined in ablation funnel
Cell carrier gas flow (l min-1) | 0.5 l min-1
ICP-MS Instrument
Make, Model and type | Thermo Element XR, Sector-field single collector ICP-MS
Sample introduction | Direct
Interface parameters | Jet interface equipped with a high-capacity pump
RF power (W) | 1525 W (tuned daily)
Make-up gas flow (l min-1) | ca. 0.98 l min-1 Ar (tuned daily) + 2 ml min-1 N2 (fixed)
Detection system | Triple (pulse counting, analog, Faraday), cross calibrated daily
Masses measured (amu) | 29, 49, 89, 91, 93, 139, 140, 141, 146, 147, 153, 157, 159, 163, 165, 166, 169, 172, 175, 178, 181, 202, 204, 206, 207, 208, 232, 235, 238
Integration time per peak/dwell times (ms) | 11 ms except at amu 49, 238 (20 ms); 206 and 207 (40 ms)
Total integration time per output data point (s) | 0.598 s
Dead time (ns) | 22
Detection efficiency | 3% U (measured with static spot, 43 mm and 5 Hz on NIST SRM612)
Typical oxide rate (248ThO+/232Th+) | 0.1%
Typical doubly charged rate (Ba++/Ba+) | 3.5%
Data Processing
Gas blank | 25 s
Calibration strategy | GJ-1 used as primary reference material for U-Pb dating; 91500, Plešovice and AUSZ7-1 as secondary (validation) reference materials. NIST SRM612 used as primary reference material for trace element quantification
Internal standard for trace element quantification | Stoechiometric Si content of zircon (15.2 wt.%); Ti concentrations corrected from matrix effects by normalization to concentration in 91500 determined by ID-ICP-MS (Szymanowski et al., 2018)
Reference Material info | GJ-1 206Pb/238U = 0.097877 ± 0.000068 ; 207Pb/206Pb = 0.060171 ± 0.00005 (Horstwood et al., 2016)
91500 (Widenbeck et al., 1995)
AUSZ7-1 (Kennedy et al., 2014)
Plešovice (Sláma et al., 2008)
NIST SRM612 (Jochum et al., 2011)
Data processing package used / Correction for LIEF | Iolite v2.5 (Hellstrom et al., 2008) with VisualAge data reduction scheme (Petrus and Kamber, 2012); LIEF corrected based on primary reference material and extrapolated to all unknowns (Paton et al., 2010)
Mass discrimination / drift correction | Normalization to primary reference material by conventional standard-sample bracketing
Common-Pb correction, composition and uncertainty | No common Pb correction applied
Uncertainty level and propagation | Uncertainty on ratios and ages are at 2σ absolute, including propagated uncertainty based on the scatter of the primary reference material (Paton et al., 2010).
Quality control / validation | AUSZ7-1 - weighted mean 206Pb/238U date = 38.27 ± 0.36 Ma (2σ, n = 15/15; MSWD = 1.3)
Plešovice - weighted mean 206Pb/238U date = 338.5 ± 1.9 Ma (2σ, n = 15/15; MSWD = 0.92)
91500 - weighted mean 206Pb/238U date = 1062.7 ± 5.6 Ma (2σ, n = 15/15; MSWD = 0.45) ; weighted mean 207Pb/206Pb date = 1075.5 ± 13.3 Ma (2σ, n = 15/15; MSWD = 0.50)
Hellstrom, J.C., et al. (2008) Iolite: Software for spatially resolved LA-(quad and MC) ICP-MS analysis, in Laser ablation ICP-MS in the Earth sciences: Current practices and outstanding issues, P. Sylvester, Editor. 2008, Mineralogical Association of Canada: Vancouver. p. 343-348.
Horstwood, M.S.A., et al. (2016) Community‐Derived Standards for LA‐ICP‐MS U‐(Th‐)Pb Geochronology – Uncertainty Propagation, Age Interpretation and Data Reporting. Geostandards and Geoanalytical Research 40(3), 311-332.
Jochum, K.P., et al. (2011). Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397–429.
Kennedy, A.K., et al. (2014) Eocene zircon reference material for microanalysis of U-Th-Pb isotopes and trace elements. The Canadian Mineralogist 52(3), 409-421.
Paton, C., et al. (2010) Improved laser ablation U-Pb zircon geochronology through robust downhole fractionation correction. Geochemistry, Geophysics, Geosystems 11, 1-36.
Petrus, J.A. and B.S. Kamber (2012) VizualAge: A Novel Approach to Laser Ablation ICP-MS U-Pb Geochronology Data Reduction. Geostandards and Geoanalytical Research 36(3), 247-270.
Sláma, J., et al. (2008) Plešovice zircon - A new natural reference material for U-Pb and Hf isotopic microanalysis. Chemical Geology 249, 1-35.
Szymanowski, D., et al. (2018) Isotope-dilution anchoring of zircon reference materials for accurate Ti-in-zircon thermometry. Chemical Geology 481, 146-154.
Wiedenbeck, M., et al. (1995) Three natural zircon standards for U-Th-Pb, Lu-Hf, trace element and REE analyses. Geostandards Newsletter 19, 1-23.