Table S2

Summary of U–Th–Pb isotope analysis by LA-ICP-MS/MS

| Laboratory and sample preparation |  |
| --- | --- |
| Laboratory name | Department of Earth Sciences, National Taiwan Normal University |
| Sample type/ mineral | Zircons from syenite |
| Sample preparation | Crushing and heavy liquid mineral separation, 1-inch resin mount, 1 μm polish to finish |
| Imaging | CL imaging by Gatan mini-CL attached to JEOL W-low vacuum scanning electron microscope, 15 kV at 0.5 nA, 1520 mm working distance |
| Laser ablation system |  |
| Make, model and type | Teledyne CETAC Analyte Excite Plus |
| Ablation cell | HelEx II Active 2-Volume Ablation Cell, 100 mm × 100 mm sample holder (25 mm in diameter × 9 pucks and 12.5 mm in diameter × 4 pucks) |
| Laser wavelength (nm) | 193 nm (ArF excimer) |
| Pulse width (ns) | < 4 ns |
| Fluence (J/cm2) | 2.00 J/cm2 |
| Repetition rate (Hz) | 17 Hz |
| Ablation duration (s) | ~24 s |
| Ablation pit depth/ ablation rate | ~20 μm pit depth, calculated assuming an ablation rate of a few dozen nm per pulse |
| Spot diameter (μm) | 25 μm |
| Sampling mode/ pattern | Static spot ablation |
| Carrier gas | 100% He in the cell; Ar make-up gas added via a co-axial mixing adapter, where the Ar gas is mixed on-axis with the He and ablated sample material, directly connected to the ICP torch |
| Cell carrier gas flow (l/min) | 0.5 l/min for MFC1 (Cell body), 0.5 l/min for MFC2 (Cup) |
| ICP-MS instrument |  |
| Make, model and type | Agilent 8900 Triple Quadrupole ICP-MS (ICP-MS/MS) |
| Sample introduction | Ablated sample aerosol via conventional tubing and the co-axial mixing adapter |
| Scan mode | MS/MS mode |
| RF power (W) | 1300 W |
| Make-up gas flow (l/min) | Sourced from Agilent 8900 ICP-MS/MS, 0.75 l/min |
| Cell gas flow rate (ml/min) | 1 ml/min for 10% NH3 + 90% He reaction gas (“10%” on 3rd mass flow controller of Agilent 8900 ICP-MS/MS), 1.0 ml/min for He collision gas |
| Detection system | Electron multiplier |
| Masses measured (Quadrupole-1 → Quadrupole-2) | 202 → 202, 204 → 204, 206 → 206, 207 → 207, 208 → 208, 232 → 232, 232 → 247, 238 → 238 and 238 → 253 |
| Dwell times (ms) | 202Hg [40 ms], 204Pb [50 ms], 206Pb [25 ms], 207Pb [35 ms], 208Pb [10 ms], 232Th [10 ms], 247Th-NH [10 ms], 238U [15 ms] and 253U-NH [15 ms] |
| Total integration time per output datapoint (s) | 0.2440 s |
| ‘Sensitivity’ as useful yield (%, element) | 0.023 % U, not including a series of reaction products of U with NH3 reaction gas such as U-NH, (Schaltegger et al., 2015) |
| IC Dead time (ns) | 30 nsec |
| Data processing |  |
| Integration time for gas blank and sample (s) | ~25 s for gas blank, ~14 s for sample |
| Calibration strategy | 91500 zircon and NIST 610 glass reference material used as primary for an external calibration of U/Pb and Pb/Pb isotope ratio, respectively. Plešovice and AS3 zircon reference materials used as evaluations. |
| Reference material information | 91500 (Wiedenbeck et al., 1995), NIST 610 (Jochum et al., 2005; Data of Jochum and Nehring 11/2006 on GeoReM database), Plešovice (Sláma et al., 2008), AS3 (Paces & Miller, 1993) |
| Data processing package used/ Correction for LIEF | Agilent MassHunter; In-house excel spreadsheet for data normalization, common Pb correction, uncertainty propagation and age calculation (Hattori et al., 2017; Sakata et al., 2014). LIEF correction assumes reference material and samples behave identically. |
| Common-Pb correction and composition | Using the Hg-corrected 204Pb, a correction for common Pb was made on the basis of the 204Pb method (Stern, 1997) and the model for the common Pb compositions (Stacey & Kramers, 1975). |
| Uncertainty level and propagation | Ages are quoted at 2σ absolute. Uncertainties from the isotopic ratio measurements, counting statistics for each analyte and common-Pb corrections were propagated. |
| Quality control/ validation | Plešovice; a weighted average of 206Pb/238U age = 333.2 ± 4.7 (2σ, MSWD = 1.3, n= 4), AS3; a weighted average of 207Pb/207Pb age = 1094 ± 39 (2σ, MSWD = 0.59, n = 4). |
| Other information |  |
|  | Sample tubing of ~3 m from the sample cell to a ICP torch. Wait time of 30 s between analyses. Washout time of c. 210 ms at 5% of the maximum peak intensity. |

References

Hattori, K., Sakata, S., Tanaka, M., Orihashi, Y., & Hirata, T. (2017). U–Pb age determination for zircons using laser ablation-ICP-mass spectrometry equipped with six multiple-ion counting detectors. Journal of Analytical Atomic Spectrometry, 32, 88–95

Jochum, K. P., Nohl, U., Herwig, K., Lammel, E., Stoll, B., & Hofmann, A. W. (2005). GeoReM: a new geochemical database for reference materials and isotopic standards. Geostandards and Geoanalytical Research, 29, 333–338.

Paces, J. B., & Miller, J. D. Jr. (1993). Precise U–Pb ages of Duluth complex and related mafic intrusions, Northeastern Minnesota: geochronological insights to physical, petrogenic, paleomagnetic, and Tectonomagmatic processes associated with the 1.1 Ga midcontinent rift system. Journal of Geophysical Research, 98, 13997-14013.

Schaltegger, U., Schmitt, A. K., & Horstwood, M. S. A. (2015). U-Th-Pb zircon geochronology by ID-TIMS, SIMS, and laser ablation ICP-MS: recipes, interpretations, and opportunities. Chemical Geology, 402, 89-110.

Sakata, S., Hattori, K., Iwano, H., Yokoyama, T. D., Danhara, T., & Hirata, T. (2014). Determination of U–Pb ages for young zircons using laser ablation-ICP mass spectrometry coupled with an ion detection attenuator device. Geostandards and Geoanalytical Research, 38, 409–420.

Sláma, J., Košler, J., Condon, D. J., Crowley, J. L., Gerdes, A., Hanchar, J. M., Horstwood, M. S. A., Morris, G. A., Nasdala, L., Norbeg, N., Schaltegger, U., Schoene, B., Tubrett, M. N., & Whitehouse, M. J. (2008). Plešovice zircon - A new natural reference material for U-Pb and Hf isotopic microanalysis. Chemical Geology, 249, 1-35.

Stacey, J. S., & Kramers, J. D. (1975). Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters, 26, 207-221.

Stern, R. A. (1997). The GSC sensitive high ion microprobe (SHRIMP): analytical techniques of zircon U–Th–Pb age determinations and performance evaluation. Geological Survey of Canada Current Research, Radiogenic Age and Isotopic Studies Report, 10, 1-32.

Wiedenbeck, M., Allé, P., Corfu, F., Griffin, W. L., Meier, M., Oberli, F., Quadt, A. V., Roddick, J. C., & Spiegel, W. (1995). Three Natural Zircon Standards for U-Th-Pb, Lu-Hf, Trace element and REE Analyses. Geostandards Newsletter, 19, 123.